Scope:

The OSBI is seeking proposals for replacement of the Deionized (DI) Water System located in the Forensic Science Center (FSC) located at 800 E. 2nd Street, Edmond, OK  73034.  Vendor will take all necessary steps to protect all property in and around the construction area.  Vendor shall remove and dispose of all debris resulting from this project offsite.  Use of OSBI receptacles and/or dumpsters are prohibited.

Project demolition consists of:

- Removal of existing ELGA Centra-R 200 DI Water unit

- Removal of DI water Mixed Bed Tank System

- Removal of existing DI Water Carbon Tank

- Removal of DI Water Softener System

Project Installation:

- Installation of DI water system capable of the following:

  - Delivering up to 30 liters per minute of type I, type II, and type III water from a distribution loop.

  - Equipped with a reservoir capacity greater than 250 liters.

  - Supply up to 18.2 MΩ.cm, type I, type II, and type III water suitable for general and analytical applications.

  - Bacterial specification <5 CFU/ml is maintained.

  - Particulate filtration down to 0.2 µm.

  - DI water system equipped with lead detection system.

  - Equipped with recirculating pump system.

  - DI water system equipped with optional Building Management System (BMS) connectivity.

- Installation of Mixed Bed Tank Control System.

  - The system will have two (2) W5TDIMB10120FSP (or equivalent) SDI Mixbed1 1.2CF FBGLS STD POT tanks.

- Installation of DI Water Carbon Tank.

  - The system will have one (1) W5TDIMRAC0120FSP (or equivalent) SDI REACT CARBON 1.2CF FBGLS STD PO tank.

- Installation of DI Water Softener System.

  - A salt-based water softener with digital controls to control the regeneration cycles is mandated.  An appropriately sized brine tank will be installed adjacent to the water softener unit to provide regeneration of the mixed bed tank.

Project Installation consists of:

Project Completion:

Other Documents:

- Bonds will be required if project exceeds $100,000.00.

- Workers COMP and General Liability insurance will be required.

Background Checks:

The OSBI shall conduct background investigations and national criminal history record checks on companies and individuals with which it contracts to provide services. The successful vendor shall submit data as required by the OSBI for each proposed employee. No employee of the vendor’s shall work on the premises until the individual is approved by the OSBI. Any person found not acceptable to the OSBI will not be allowed to perform work at any OSBI facility.